## Slide 1: Surfaces, Interfaces & Beyond
ELECTRON SPECTROSCOPIES FOR SURFACE ANALYSIS:
X-ray Photoelectron Spectroscopy
Auger Electron Spectroscopy
UW- Madison Geoscience 777
Peter Sobol
Department of Geoscience
University of Wisconsin - Madison
March 12, 2019

## Slide 2: What’s the point?
UW- Madison Geoscience 777
With SEMs and electron microprobes we image materials and determine chemistry using electrons to produce secondary and backscattered electrons and X-rays – giving us certain information – based upon the “interaction volumes”of 10-20 keV electrons
There are other valuable related techniques, with some similarities, also using electrons and X-rays, which give additional information, typically of much shallower regions of sample: surface techniques:

## Slide 3: 4 Analytical techniques traditionally classified as “Surface” techniques
X-ray Photo-electron Spectroscopy (XPS) –
  Excitation: “Monochromatic” Soft X-rays (<1.4keV)
  Analyzed species: Inelastic “photo”-electron electrons
Auger Electron Spectroscopy (AES) –
  Excitation: High energy electrons (3-15keV)
  Analyzed species: Inelastic “Auger” secondary electrons
Secondary Ion Mass Spectrometry (SIMS)–
  Excitation: High energy ions (5-10keV)
  Analyzed species: Secondary ions.
Low Energy Ion Scattering(LEIS,ISS)
  Excitation: low energy ions <500eV
  Analyzed species: Scattered primary ions
UW- Madison Geoscience 777

## Slide 4: “Surface Analysis” Definitions
Sensitive to the volume of a material adjacent to an interface where the properties (structure, composition, chemistry) are influenced by the presence of the interface, as opposed to the bulk.
Typically the outer (1-20) atomic layers: depths of 0-100 angstroms. Sub mono-layer sensitivity.
Why is this important?
Most interactions between materials are determined by the surface properties, not the bulk properties of the material
“Bulk” material properties are often a function of internal interface properties.
UW- Madison Geoscience 777

## Slide 5: Interfaces and material properties:
UW- Madison Geoscience 777

## Slide 6: Surface Sensitive Techniques
Depth of Analysis
Courtesy of EAG Laboratories
UW- Madison Geoscience 777

## Slide 7: Surface Sensitive Techniques
Courtesy of EAG Laboratories
UW- Madison Geoscience 777

## Slide 8: AES application: Pre-solar grains
https://journals.uair.arizona.edu/index.php/maps/article/viewFile/15752/15740
SEM and Auger maps elemental maps
Pseudo-color overlay (upper right) identifies unique material phases
Magenta phase is identified as a forsteritic olivine (fosterite) phase ~300nm in size.
Smaller grains of similar composition are visible in the image.
UW- Madison Geoscience 777
Just one example showing what is possible:

## Slide 9: A little History
XPS:
1887 Heinrich Hertz observes metals illuminated with UV light emit electrons- but with a low frequency cutoff, regardless of intensity
1905 Einstein explanation: light is quantized (photons) each with energy hʋ. important step in development of quantum theory – (Nobel prize)
1954- Kai Siegbhan at University of Uppsula, Sweden develops technique and instrumentation. 1967 publishes reference book: “Electron Spectroscopy for Chemical Analysis”: ESCA – (Nobel prize)
AES
Auger process discovered in 1923 by Lise Meitner and independently in 1925 by Pierre Auger
1953 Lander proposes using Auger electrons for analysis
1960’s Laboratory instruments. ~1970 first commercial instruments (Physical Electronics, Inc., Eden Prairie MN).
UW- Madison Geoscience 777

## Slide 10: Comparison of Surface Analysis Techniques
AES XPS SIMS EDS/WDS
Probe Beam Electrons X-rays Ions Electrons
Analysis Beam Electrons Electrons Ions X-rays
Spatial Resolution 0.008 µm 10 µm (1mm) 0.07 µm 0.5-3 µm
Sampling Depth(nm) 1-10 1-10 1-2 300-3,000
Detection Limits 0.1atom % 0.01atom % ppm 0.1 atomic%
Information Content Elemental Elemental Elemental Elemental
Chemical Chemical Molecular
Quantification Fair Excellent Std. needed Good
UW- Madison Geoscience 777

## Slide 11: But: Surface Analytical techniques are frequently extended and/or extrapolated in order to characterize bulk material properties!
UW- Madison Geoscience 777

## Slide 12: X-ray Photoelectron Process
2p
1s
2s
1/2
3/2
Efermi level
Binding Energy
Photoelectron
Binding Energy = X-ray Energy (hν) - Photoelectron Kinetic Energy
X-ray
Evacuum level
UW- Madison Geoscience 777

## Slide 13: XPS: Photon-electron interactions:
Photon interactions with charged particles are “all or nothing”: a photon normally gives up all its energy or doesn’t interact at all (required to satisfy both conservation of momentum and energy!)
Photoelectron Kinetic Energy = hν – Binding Energy*
Photons may interact with any electron where E(photon) > Binding Energy.
All elements have at least one electron orbital with binding energy <1.1KeV
* there’s always an asterix!
UW- Madison Geoscience 777

## Slide 14: XPS: Au Survey spectrum
All occupied orbitals with E<hv (photon energy) will produce photoelectrons.
Area ratios between different orbital peaks will be consistent (aid to identification)
Orbitals with angular moment l > 0 (everything but s orbitals) exhibit “Spin-orbit” splitting into 2 lines with fixed ratios.
  …
Valence Band
UW- Madison Geoscience 777

## Slide 15: Auger Electron Process
2p
1s
2s
1/2
3/2
Efermi level
Binding Energy
Auger Electron
Starts with a vacancy in a core orbital
Process involves 3 electrons (usually denoted in spectroscopic notation)
Auger Electron Kinetic Energy:
EKLL = BEK – BEL – BEL
(binding energy for orbitals in the presence of core hole, not ground state of atom!)
Independent of origin of core hole!
Core Vacancy (created X-ray, electron, or ion)
K
LI
LII
LIII
Kinetic Energy
Evacuum level
UW- Madison Geoscience 777

## Slide 16: Auger: Displaying data
Auger lines typical consist of broad peaks on a large curved secondary background: E N(E)
 Commonly displayed as differentiated data: E dN(E)/dE
Peak positions are usually listed as the position of the negative going peak in the differentiated data.
Because there are 3 electrons involved in the transition – and many possible combinations, spectra can have many peaks, especially for heavier elements .
0
500
1000
1500
2000
2500
3000
Kinetic Energy (eV)
E N(E)
E N(E) x 5
EdN(E)/dE
Cu LMM
Cu MNN
UW- Madison Geoscience 777

## Slide 17: Auger: Spectra
Cu MNN Line ~60eV
Cu LMM series from 700-900eV
All elements have major lines <2keV
UW- Madison Geoscience 777

## Slide 18: Auger: Spectra and images
Auger is inherently an imaging technique performed with an electron beam similar to EPMA
By plotting peak intensity vs. position we can create elemental and chemical maps of a surface.
.
UW- Madison Geoscience 777

## Slide 19: Auger Electron vs. X-ray Emission
Incident
Beam
X-ray
Photon
X-ray Fluorescence
(detected in EDS/WDS)
Auger Electron
X-ray
X-ray fluorescence and Auger Electron emission are competing relaxation processes of a core electronic vacancies.
Auger process dominates for low atomic number elements!
Auger Electron Emission (detected in AES)
or
UW- Madison Geoscience 777

## Slide 20: Correcting for the work function
The difference between the fermi level and the vacuum level (0eV potential) is the “work function”. Equal to the energy required to extract the least bound electron.
When in electrical contact the fermi levels of dissimilar materials equalize (otherwise current would flow!)
Need to adjust the measured Kinetic energy by the work function of the spectrometer for both Auger and XPS.
Usually a calibration parameter of the instrument
What about insulators?!
UW- Madison Geoscience 777

## Slide 21: Correcting for surface charge
Insulating materials have surfaces that are not in electrical contact with the spectrometer!
Surface charge may accumulate – shifting the kinetic energy of outgoing electrons.
X-rays don’t impart a charge to a surface, but secondary electrons leaving do. Charge can generally be balanced by providing a low energy flood of electrons 0-2eV to the surface – they will be attracted to areas that have charged positive and “neutralize” the charge.
  Generally a stable condition can be achieved, but with a small net charge
Auger uses an electron beam for excitation, charging is generally not controllable on good insulators, and the technique not be used.
UW- Madison Geoscience 777

## Slide 22: Correcting for surface charge
  Need to have a way to correct the energy scale for any residual surface charge after neutralization.
  Nature provides: virtually all surfaces exposed to the atmosphere will have a small amount of “adventitious” carbon (short chain C-C,C-H bonds).
  Generally agreed to appear at ~284.8eV in XPS.
  Shift = Observed C1s peak -284.8
  Shift ALL spectra the same amount
UW- Madison Geoscience 777

## Slide 23: Analysis Depth
Electrons in a solid can only travel a short distance before interacting with other electrons or atoms.
Most such interactions are “elastic” – energy is transferred from the electron.
Only electrons emitted within the first several atomic layers have a significant chance of escaping the surface with their original energy.
Most others will be absorbed, but some will be “scattered” and escape the surface, but with lower energy.
Electrons emitted near the surface can escape at any angle, but electrons from deeper layers need to be travelling nearly perpendicular.
Photoelectrons and Auger electrons are characterized by their energy – if they elastically scatter they are lost to the analysis
UW- Madison Geoscience 777

## Slide 24: The convenient Electron Volt:
38 meV: average kinetic energy of a thermal air molecule at STP.
720 meV: the energy required to break a covalent bond in germanium
1.1 eV: the energy required to break a covalent bond in silicon
1.6 eV to 3.4 eV: the photon energy of visible light
0.1 eV to 10 eV: per bond energy of molecular bonds
13.6 eV: the energy required to ionize atomic hydrogen;
10-1100eV: all elements have some core level orbitals in this range
1486eV: Energy of the Aluminum K-alpha X-ray photon.
1 MeV (1.602×10−13 J): about twice the rest energy of an electron
200 MeV: the average energy released in nuclear fission of one U-235 atom
Unit of energy defined as the amount of work required to move a particle with a charge of 1e ‘uphill’ through a potential difference of 1 volt. Conversely it is equal to the kinetic energy gained by an electron when it is accelerated ‘downhill’ through potential difference of 1 volt. 1 eV x 1 Coulomb = 1 Joule
An electron column might have its source potential at -10kV – when an electron from this source reaches a ground potential, it will have 10keV of kinetic energy.
UW- Madison Geoscience 777

## Slide 25: Analysis Depth
Inelastic Mean free path (λ)is ~1-5nm (10-50A) for typical XPS and AES measurements
Beer’s law: I = I0 exp(z/ λ), For z= 3 λ, I/ I0 = 0.05. 95% of signal comes from < 3 λ
UW- Madison Geoscience 777

## Slide 26: EDS vs. Auger
“Spot size” determined by incident beam. - With Field Emission Electron Source minimum lateral resolution <10 nm - (EDS with the same source: analysis area > 1um)
Analysis depth <10nm
Nano-volume analysis capabilities.
Note: Secondary and backscattered electrons will be detected along with Auger electrons.
Secondary Electrons
Backscattered Electrons
UW- Madison Geoscience 777

## Slide 27: Instrumentation
Sample
Electrons
Excitation Source X-rays
Electrons
Electron Energy
Analyzer & Detector
UHV Vacuum Envelope
Instrumental essential features:
Instrumental essential features:
  Vacuum envelope (and pumping systems
  Excitation source (XPS:X-rays, AES: Electrons)
  Electron Energy Analyzer and detector
Additional features:
  Charge neutralization sources (low energy electrons, ions)
  Sputter ion guns (surface cleaning, depth profiling)
UW- Madison Geoscience 777

## Slide 28: Vacuum requirements
For a gas, the mean free path (distance a particle can travel without a collision):
  p=pressure, d=Molecular diameter
Electrons scattered between the sample and detector reduce signal and increase background noise.
For practical analyses, I >> path length from sample to detector (~1-2m).
Pressure < 1e-4 Pa (~5e-7 torr)
UW- Madison Geoscience 777

## Slide 29: Vacuum requirements
At a surface exposed to a gas, particles impact the surface at a rate proportional to pressure. A high percentage of those particles will “stick” to most clean surfaces.
Particle flux on a surface:
  particle density. = average velocity
From ideal gas law:
UW- Madison Geoscience 777

## Slide 30: Vacuum requirements
Monolayer formation time is equal the number of sites on a surface divided by the impingement rate (flux) and the stiction coefficient (probability a particle remains on the surface)
, S = stiction coefficient =1
  A typical surface may have 1015 sites/cm2
  At 1 atm = 3ns,
  at 0.1pa, =3e-3 s
  at 10-4 pa, = 3s,
  at 10-7 pa, = 3000s
Pressures < 1e-6 pa (UHV range) required to maintain surface cleanliness during analysis.
UW- Madison Geoscience 777

## Slide 31: Vacuum requirements
Residual gas species in a vacuum chamber consist primarily of H, H20 (-> OH, O), CH4, N, CO and organic compounds
Most are likely to react with surfaces on contact, especially in the reducing environment created by energetic electrons
“stiction” coefficient ~1
UW- Madison Geoscience 777

## Slide 32
Characteristics of an X-ray source for XPS:
Approximately Monochromatic for narrow distribution of photo-electrons
Appropriate energy: 1-5keV
Low background electrons, x-rays, heat
“Standard” X-ray source
“Standard” Source
Electrons bombard metal surface with high energy (10-15keV)
X-rays produced by fluorescence flood sample
Typically multiple X-ray energies
Bright, simple, low cost
1/r2 loss of intensity (must be close to sample)
Sample exposed to heat, electrons excited from window.
Satellite X-ray peaks.
Instrumentation: X-ray Sources (XPS)
UW- Madison Geoscience 777

## Slide 33: Instrumentation: X-ray Sources (XPS)
To overcome limitations of standard source, most instruments are now equipped with an X-ray Monochromator:
Aluminum K-alpha electron has wavelength ~ lattice separation in quartz
  Quartz crystals can be used to diffract Al Kα X-rays through reasonable geometries.
  By placing the quartz crystals on an appropriately curved substrate, X-rays can be micro focused onto sample surface ( <10um spot sizes). Brightness limit is the flux density of electrons the Anode can withstand.
  Scanning electron beam on anode scans X-ray beam on sample
  ~0.2eV linewidths.
Low backgrounds, no exposure to heat/electrons from source
More complicated and expensive
UW- Madison Geoscience 777

## Slide 34: Instrumentation: X-ray Sources (XPS)
Synchrotron sources
  Extremely high brightness
  Micro focusing
  Tunable X-rays
  Not laboratory Scale
UW- Madison Geoscience 777

## Slide 35: Instrumentation: AES sources
Focused Electron beam sources similar to those found in SEM & EPMA instruments
  Tungsten filament
  LaB6
  Field Emission
But sources typically optimized for higher brightness rather than lateral resolution to provide better S/N for low intensity Auger Electron signals.
  3nA – 1uA currents
UW- Madison Geoscience 777

## Slide 36: Instrumentation: Electron Analyzers
Requirements for XPS:
  Need to resolve peaks of a few 10ths of an eV  High energy resolution
  Signals can be small  adjustable transmission
  Unfocused sources  analysis are defined by analyzer
  Emission angle contains information  Adjustable angular acceptance (high transmission vs. narrow acceptance)
Requirements for AES:
  Auger signals have poor S/N High transmission
  Auger signals tend to be wide  Modest energy resolution
  Microscopically natural surfaces are often rough Insensitivity to sample geometry & shadowing
  Analysis area always defined by beam, not analyzer.
UW- Madison Geoscience 777

## Slide 37: Instrumentation: XPS Electron Analyzers
Most XPS systems use a Hemispherical Capacitor Analyzer with input lenses:
  Hemisphere can be adjusted for high transmission, or high energy resolution (0.1%/E)
  Input lens provides high flexibility:
    Adjustable area definition & angular acceptance
    Adjustable energy retardation to optimize resolution vs. transmission
  The retarding lens determines the kinetic energy of the electrons which enter the analyzer. “Scanning” over an energy range is performed by adjusting the retarding lens
  The difference between the inner and outer spheres determines the “pass energy” window: Transmission vs. energy resolution.
    Normally run in “Fixed Analyzer Transmission” mode: energy resolution is the same across the energy spectrum
  To improve sensitivity a position sensitive detector oriented along the energy dispersion axis of the analyzer allows electrons of a range of energies to be counted in parallel, while retaining energy resolution.
UW- Madison Geoscience 777

## Slide 38: Instrumentation: XPS Electron Analyzers
  Small area XPS analysis area can be performed in two different ways.
    “Zoom lens” only accepts electrons from a small area
    Focused source: only excites a small area of the surface.
UW- Madison Geoscience 777

## Slide 39: Instrumentation: AES Electron Analyzers
AES systems frequently use a Cylindrical Mirror Analyzer:
Very high transmission: wide acceptance angle, large energy pass window
Electrons can be emitted at any azimuthal angle and still enter analyzer
Because the electron trajectories don’t travel down the axis of the analyzer, an electron gun can be placed coaxially.
  No shadowing, insensitive to surface roughness/geometry!
Usually run in “fixed retard ratio” (FRR) mode: Transmission increases and energy resolution decreases with increasing energy.
Poor energy resolution: 5% of E.
UW- Madison Geoscience 777

## Slide 40: XPS Survey Spectra
“Survey” wide energy scan, high transmission, low resolution: polyethylene terephthalate example:
Note oxygen Auger line!
But why is the x-axis backwards!!
Note step in background at below each peak (in KE). These are electrons from the peak that are elastically scattered before leaving the sample: Average origin is deeper in the sample than the primary peaks.
BE = hv – KE
By measuring the area under the O1s and C1s peaks, we can quantify the elements
486 686 886 1086 1286 1486
Kinetic energy (ev)
UW- Madison Geoscience 777

## Slide 41: XPS Elemental quantification
Intensity of a photoelectron peak (cps) is given by
For a given analyzer condition, all variables except n are constant. So we can assign an atomic sensitivity factor S:
The ratio of two elements:
The atomic fraction of constituent x:
Values of S are tabulated by instrument manufacturers (and embedded in instrument software).
UW- Madison Geoscience 777

## Slide 42: Auger Elemental quantification
Ditto what I just said about XPS
Auger peak intensities typically measured by peak-to-peak height of differentiated spectrum
Sensitivity factors are much more dependent on chemical state due to complexity of line shapes, variable backgrounds and differentiation
Sensitivity decreases for heavier elements.
UW- Madison Geoscience 777

## Slide 43: XPS Quantification
Elemental quantification comparable to EPMA, when materials are homogenous within the analysis volume.
Much smaller analysis volume!
Care must be taken in the presence of interface effects: contamination, oxidation, relaxation
https://www.researchgate.net/publication/230348878_Quantitative_XPS_Measurements_of_Some_Oxides_Sulphides_and_Complex_Minerals
UW- Madison Geoscience 777

## Slide 44: XPS: Clues to the Chemistry
Initial state effects – “It was like that when I got here”
  Chemical shift: Bonding environment of atom: increases or decreases actual binding energy of electron
  Line widths: – width of line is dependent on chemical state
    Lifetime of core vacancy (ΔEΔt = ħ)
    Vibrational-rotational broadening
Final state effects – Outgoing electron excites additional discrete transition within atom or material.
Always reduces apparent binding energy
  Asymmetry
  Shake-up & Shake-off:
  Plasmons
Virtually always related to “nearest-neighbor” interactions.
Powerful tool for
UW- Madison Geoscience 777

## Slide 45: XPS: Initial state: Chemical shifts
Electron orbitals are ‘fuzzy’ probability distributions surrounding the nucleus.
To the extent that valence orbitals overlap core orbitals, they ‘screen’ the core orbital from the charge of the nucleus, reducing the binding energy of the core orbital electron.
As valence electrons form chemical bonds with other atoms, they are pulled into different orbitals that have more or less overlap with the core orbitals, resulting in a shift of the core orbital binding energy.
Oxidation: Cation core orbitals shift to lower binding energy, while the anion core orbitals shifts to higher binding energy
UW- Madison Geoscience 777

## Slide 46: XPS: Initial state: Chemical shifts
High resolution scan of PET reveals structure of C1s peak, revealing chemical “shifts” due to different bonding environments within the PET molecule.
Quantification also works for when comparing chemical states.
No sensitivity corrections required- core level photoelectron cross sections are insensitive to bonding environment
UW- Madison Geoscience 777

## Slide 47: XPS: optimizing analyzer resolution
Decreasing analyzer resolution increases sensitivity.
Optimize instrument settings for best compromise depending on resolution and sensitivity required.
UW- Madison Geoscience 777

## Slide 48: XPS: Clues to the chemistry
Magnitudes of chemical shifts for large numbers of compounds are found in the literature and tabulated in the “Handbook of X-ray Photoelectron Spectroscopy” and various on-line databases.
UW- Madison Geoscience 777

## Slide 49: XPS Final State effects: Asymmetry
Conducting (and semiconducting) materials have a high density of energy states of the valence (conduction band) electrons with binding energies of zero to a few eV.
Photoelectrons have a high likelihood of interacting with these valence electrons and losing a little bit of energy
Results in an asymmetric tail to the low BE side of the peak.
Seen in virtually all conducting materials
Asymmetric Tail on C1s in Graphite
UW- Madison Geoscience 777

## Slide 50: XPS Final state effects: Shake-up
Outgoing photo-electron may excite other discrete transitions.
Graphite is a conductor with covalent bonds shared around 6 membered rings of carbon
These “π” orbitals have two possible configurations: bonding and anti-bonding. The outgoing electron can ‘flip’ these from bonding to anti-bonding, giving up ~6eV in the process.
Characteristic of “benzene ring” structure in organic molecules
π  π* shake-up satellite
UW- Madison Geoscience 777

## Slide 51: Final state Effects: Shake-up
In transition metals and rare-earths, strong shake-up lines typically occur in paramagnetic states (outer shell unpaired electrons)
Absence of shake-up indicates elemental or diamagnetic states
  Often useful in determining chemical state where observed.
UW- Madison Geoscience 777

## Slide 52: XPS Final state effects: Plasmons
A “Plasmon” is an oscillation mode of the free electron plasma found in the repeating lattice of a conducting material.
Because the plasma is ubiquitous there can be a high probability of interaction with an outgoing photoelectron  increasing the apparent binding energy.
The photoelectron can react more than once, resulting in a series of peaks to the higher energy side.
Conductors have two possible plasmon energies, 2d and 3d that differ in energy by
Most strongly observed in aluminum and silicon.
3d plasmon sat.
2d plasmon sat.
2x
3x
UW- Madison Geoscience 777

## Slide 53: Auger Parameter
Cu2O very weakly distinguished from Cu metal by either lineshift or shape, but in the KLL Auger peak the shift is larger than CuO!
UW- Madison Geoscience 777

## Slide 54: Auger parameter
Define the “Auger parameter” as the difference between an Auger and a photoelectron line.
Substituting:
This expression allows us to tabulate Auger parameter values that are independent of photon energy.
UW- Madison Geoscience 777

## Slide 55: Auger parameter values – independent of photon energy:
UW- Madison Geoscience 777

## Slide 56: XPS Chemical State Imaging
500 µm FOV – 10 µm x-ray beam
96
98
100
102
104
106
108
Binding Energy (eV)
Si Oxide
Si Oxy-Nitride
Silicide
WxSiyOz
8 Min. Elemental Images
Si Chemical State Images
Extracted Si Basis Spectra
128 channel Si basis spectra used to create chemical state maps with LLS Fitting
Si oxide (green), Si oxy-nitride (blue), W silicate/silicide (red)
N 1s
O 1s
W 4f
Si 2p
Colored dots indicate the locations basis spectra were extracted from
500 µm FOV – 10 µm x-ray beam
Oxy-nitride
Oxide
Silicide
Most recent generation of XPS instrument include the ability to scan the X-ray monochromator, or the acceptance area of the analyzer to create images
Provides imaging down to ~10um, but with low S/N
UW- Madison Geoscience 777

## Slide 57: Chemistry by Auger
Auger electrons are also sensitive to initial and final state effects.
3-electron transitions create more complicated response to chemical states
Chemical state interpretation possible through peak shift and line shape analysis.
More reliant to comparison to reference materials.
Linear least squares fitting, Principal Component Analysis and Target Factor Analysis are some of the numerical techniques used to identify and quantify chemical state analysis of Auger spectra.
UW- Madison Geoscience 777

## Slide 58: Auger Chemical State Analysis
High Energy Resolution Imaging of Semiconductor Chemical States
A
B
C
Composite Si KLL
Elemental Si
Silicide
Si Oxynitride
Si KLL image with ROI areas
Si KLL Basis Spectra
1604
1608
1612
1616
1620
1624
Kinetic Energy (eV)
Intensity
1605
1615
1625
Kinetic Energy (eV)
Intensity
LLS fitting from the basis spectra extracted from regions of interest used to create chemical state maps of Si with 50nm resolution
Silicide
Si Metal
200 µm FOV SEI of a semiconductor bond pad
0.1% energy resolution data
A
B
E
D
SEI
Elemental Si
Si Oxynitride
Silicide
C
Si KLL Peak Area
F
Si Chemical States
UW- Madison Geoscience 777

## Slide 59: Auger Chemical State Analysis
High Energy and Spatial Resolution Imaging of Semiconductor Chemical States
200 µm FOV SEI of a semiconductor bond pad
0.1% energy resolution data
A
B
E
D
SEI
Elemental Si
Si Oxynitride
Silicide
C
Si KLL Peak Area
F
Si Chemical States
Auger Color Overlay 20kV 10nA 256x256 pixels
Green: Elemental Si Blue: Si Oxynitride Red: Silicide
Silicide (red)
Oxynitride (blue)
Elemental Si (green)
UW- Madison Geoscience 777

## Slide 60: XPS Application: Alteration of olivine
Alteration of Olivine in a Carbonate rich aqueous environment at 70C
Initially surface shows:
Mg:Si:0 1:1:2 +1%Fe
Fe:2+
Si: silicate
C: Hydrocarbon
After reaction:
Si: SiO2
Mg:2+
O: SiO2
Carbonate
UW- Madison Geoscience 777

## Slide 61: XPS Applications: Kinetics of radiation induced changes in olivine
Simulates radiation exposure in space
Reference spectra of Fe in Fe0, Fe2+ and Fe3+ states
Exposure to low fluences of He+ ions shifts the Fe2+ to Fe3+ at the surface
Exposure to higher fluences of He, or Ar ions reduces the Fe2+ to Fe0
Simultaneously the Mg/Si and O/Si ratios increase, indicating the formation of a surface layer MgxOy surface layer
UW- Madison Geoscience 777

## Slide 62: Auger Analysis of Thin Surface Defect
200
400
600
800
1000
1200
1400
1600
1800
2000
Kinetic Energy (eV)
c/s
C
O
Si
Cu
Cu
O
C
Point 4
Point 3
Point 1 - Defect
Point 2 - Defect
Auger spectra (30 seconds each) show that the defect is Cu
62
Neither Auger nor XPS require prior knowledge of the sample for analysis- commonly used forensic techniques for surface defects/contaminants
UW- Madison Geoscience 777

## Slide 63: AES application: Ti/ TiB2 / SiC Interface
(a) SEM image- SIC on the right, Ti on the left. TiB2 intermediate layer
(b) Boron Auger map shows interaction of Boron containing coating with Ti matrix.
Composite image Carbon (green), Titanium (blue) and Boron (red) shows light blue phase. Spectrum indicates presence of TiC in this region
UW- Madison Geoscience 777

## Slide 64: AES application: Pre-solar grains
Olivine Reference: No distortion to charging
https://journals.uair.arizona.edu/index.php/maps/article/viewFile/15752/15740
Comparison of AES and EPMA on reference samples
Pyroxene and Olivine
Can be distinguished
UW- Madison Geoscience 777

## Slide 65: AES application: Pre-solar grains
https://journals.uair.arizona.edu/index.php/maps/article/viewFile/15752/15740
SEM and Auger maps elemental maps
Pseudo-color overlay (upper right) identifies unique material phases
Magenta phase is identified as a forsteritic olivine (fosterite) phase ~300nm in size.
Smaller grains of similar composition are visible in the image.
UW- Madison Geoscience 777

## Slide 66: AES: In situ fracture analysis
In Situ fracture exposes a fresh surface of a material within the UHV analysis chamber.
Preserves the chemistry of the exposed fractured surface from reaction with air.
Common and powerful method for materials studies where grain boundary interfaces often determines bulk material properties
In this example the grain boundary can be 30+% P even though P is 0.05% of bulk composition
https://doi-org.ezproxy.library.wisc.edu/10.1002/sia.740100405
UW- Madison Geoscience 777

## Slide 67: Going beyond the surface
Surface and near surface composition & chemistry often reflects bulk properties and can be extrapolated in many circumstances.
“Inert” materials show little difference between surface and bulk: polymers, oxides, crystalline materials (excluding metals). i.e surface composition can be extrapolated to bulk composition
Varying of surface sensitivity can deconvolve surface and bulk influences (angle resolved analysis)
Ion “sputtering” can be used to expose the sub-surface of a material (Depth profiling).
Sometimes the near surface volume is all the bulk you have: thin films in semiconductors, small grains
UW- Madison Geoscience 777

## Slide 68: Going deeper, and shallower.
Changing the effective analysis depth :
XPS
“Angle Resolved” analysis – tilting sample or changing azimuthal acceptance angle. Grazing angles are more surface sensitive.
Comparing signals with different energies (hence escape depth)
Changing X-Ray excitation energy: Mg<Al<Cr<Zr or synchrotron sources
Ion Sputtering depth profiling
Auger:
“Angle Resolved” analysis
Comparing signals from different energy lines
Ion Sputtering depth profiling
K-Alpha
Photon Energies:
Mg 1253eV
Al 1486eV
Si 1739eV
Ti 4505eV
Cr 5405eV
Zr 15690eV
UW- Madison Geoscience 777

## Slide 69: Angle Resolved Analysis
Beer’s law: I = I0 exp(z/ λ).
Modified for “take off angle” θ: I = I0 exp(z cos θ/λ),
Consider a thin layer A sitting on top of bulk material B:
Signal from A (integrated from depth d to 0:
Signal from B attenuated by passing through A:
Thickness d can be determined from the measured ratio of the signals from A and B.
  Can be generalized to multiple layers.
Analysis at multiple angles can distinguish: uniform overlayers vs. islands vs. multiple layers vs. gradients.
UW- Madison Geoscience 777

## Slide 70: XPS: Angle Resolved Examples
UW- Madison Geoscience 777

## Slide 71: XPS X-ray energy profiling:
Cr K-alpha monochromator now available on some instruments (5420eV)
UW- Madison Geoscience 777

## Slide 72: XPS X- ray energy profiling
Pyrite (FeS2) fresh cleaved (001) surface in vacuum
Analysis with 1487eV X-ray shows surface contribution
Analysis with 260eV (synchrotron source) allows shows details of S shifts due to surface states
Consistent with cleavage only Fe-S bonds
UW- Madison Geoscience 777

## Slide 73: XPS X- ray energy profiling
Pyrite (FeS2) fresh cleaved (001) surface in vacuum
Analysis with 1487eV X-ray shows surface contribution
Analysis with 260eV (synchrotron source) allows
shows details of S shifts due to surface states
Consistent with cleavage only Fe-S bonds
UW- Madison Geoscience 777

## Slide 74: Sputter Depth profiling
Energetic ion bombardment can be used to “sputter” the sample surface, removing material sequentially from the surface.
Alternates bombardment with spectroscopic analysis to build up a “profile” of the structure as it is “drilled down”
“Traditionally” done with a high energy 0.5-4keV Ar+ ion beam
  Typical sputter rates: 0.5-2 nm/min
UW- Madison Geoscience 777

## Slide 75: 500 eV Ar+ Depth Profile of Low Emissivity Glass Coating
Sputter Depth profiling
0
50
100
150
200
250
300
0
10
20
30
40
50
60
70
80
90
100
Sputter Time (min)
Atomic Concentration (%)
N
Si
Al
O
Zn
Zn
O
Sn
Ag
Ag
Ti
Ti
O
Zn
Sn
Si
C
Sn
O
Al
Mg
Ca
Na
~200 nm total film thickness
Complex structures microns thick can be analyzed
UW- Madison Geoscience 777
Notes: low-e coating-02.102.Low-e Glass+1_1.pro: low-e coating- 500eV Ar+ 750nA 3x3 mm with Zalar 1 min/cycle 9A/min vs. SiO2 2016 Aug 22 Al mono 50.0 W 200.0 µ 45.0° 140.00 eV PE 9A/min vs SiO2

## Slide 76: Sputter Depth Profiling
Fe7S8 Sample proposed to have thin layer of surface oxide Fe(III)-Oxyhydroxide
Depth profile shows an intermediate Fe2S3 Layer ~5 monolayers thick
UW- Madison Geoscience 777

## Slide 77: Sputter depth profiling has a problem!
Energetic ion bombardment tends to change the sample chemistry and stoichiometry
Reducing environment: will reduce most inorganic oxides significantly,
Preferential sputtering will change measured elemental ratios.
UW- Madison Geoscience 777

## Slide 78
Sputter depth profiling:
20
10
Sample provided by Dr. Chris Szakal of NIST
0
50
100
150
200
0
20
40
60
80
Sputter Time (min)
Atomic Concentration (%)
Si 2p
C 1s
O 1s
0.25 kV Argon
Sputter Rate for SiO2 ~ 0.3 nm/min
Sputter Rate for PMMA ~ 0.4 nm/min
Expected
Concentration
Sputtering is particularly damaging to organic materials
PMMA and SiO2 at interface reduced
UW- Madison Geoscience 777

## Slide 79
Sputter depth profiling:
0
4
8
12
16
20
0
20
40
60
80
100
Sputter Time (min)
Atomic Concentration (%)
C 1s
O 1s
Si 2p
10 kV C60
Sputter Rate: SiO2 ~ 0.4 nm/min
Sputter Rate for PMMA ~ 8.3nm/min
Sample provided by Dr. Chris Szakal of NIST
C60 sputtering maintains expected chemical composition throughout the PMMA thin film
Recent developments in Ion gun technology:
C60 (fullerines) ion gun
Gas Cluster Ar+ ion guns (2500AMU)
Both source bombard the surface with “big”, slow moving molecular structures.
Only interact with outer monolayer: preserve chemistry and stoichiometry of subsurface layers.
Effective on organic and inorganic materials
UW- Madison Geoscience 777

## Slide 80: GCIB Depth Profile of Organic LED Layers
Sample courtesy of Prof. Russell Holmes, U of MN
395
400
405
410
0
20
40
60
80
Binding Energy (eV)
N1s (BPhen)
N1s (TCTA)
0
50
100
150
0
20
40
60
80
100
Sputter Depth (nm)
Atomic Concentration (%)
C 1s
Si 2p
O 1s
10 kV Ar2500+ Sputtering
N 1s (BPhen)
10X
N 1s (TCTA)
10X
Si
TCTA Gradient
BPhen Gradient
100%
0%
100 nm
Native Oxide
Sputter Depth Profile of a Graded OLED Test Structure
BPhen
TCTA
UW- Madison Geoscience 777
Notes: Tris(4-carbazoyl-9-yltriphenyl)amine
  (TCTA) (Hole Transfer Material)
  6.9 atomic % N 93.1 atomic % C
  Higher N 1s Binding Energy
Bathophenanthroline
  (BPhen) (Electron Transfer Material)
  7.7 atomic % N 92.3 atomic % C
  Lower N 1s Binding Energy

## Slide 81: AES Charge Neutralization
Cluster Ar ion sources at low energies and doses can provide effective charge neutralization during Auger analysis on “moderate” insulators
  Electrically isolated conductors
  Thin non conducting layers over conductors
  Materials with moderate electrical resistance
UW- Madison Geoscience 777

## Slide 82: Going beyond the surface
Surface electron spectroscopies are:
“Easy”:
  Insensitive to material characteristics (other than vacuum compatibility)
  Require little or no material preparation (polishing, coating, etc)
  Standards generally not required
Comprehensive:
  Agnostic to composition (sensitive to nearly all elements)
  Provide elemental and chemical identification
  Quantitative for both
UW- Madison Geoscience 777