## Slide 1: SIMS Basics

## Slide 2: Secondary Ion Mass Spectrometer (SIMS)
±10kV
Secondary Ion
Primary Ion
(+10kV/−13kV)
Magnetic Field
Electrostatic Field
Double focusing
Mass Spectrometer
18O
16O
Ion counters
Sample
E-gun
**Surface analysis**
Simplified Schematic of SIMS
Secondary ions are produced by the sputtering of primary ions
fixed magnet strength
“multi-collection mode”
Notes: -Because this is a mass spectrometer, we can distinguish and compare ratios of different isotopes of a given element or compound.
-In the sims configuration, there are two configurations to detect secondary ions.
In the first configuration the magnet has a fixed field strength. In this configuration:
-the lighter isotopes are more sharply bent in their travel path when compared to heavier isotopes. At the end of the travel path we have
multiple collectors to detect and count each isotopologue. This is referred to as a “multi-collection” configuration.

## Slide 3: Secondary Ion Mass Spectrometer (SIMS)
±10kV
Secondary Ion
Primary Ion
(+10kV/−13kV)
Magnetic Field
Electrostatic Field
Double focusing
Mass Spectrometer
18O
16O
Ion counters
Sample
E-gun
**Surface analysis**
Simplified Schematic of SIMS
Secondary ions are produced by the sputtering of primary ions
16O or 18O
variable magnet strength
“peak-jumping” or “mono-collection” mode
Notes: In a second type of collection, the magnet strength is varied, which allows isotopes of different masses to be steered onto one detector. This is referred to as either peak-jumping or mono-collection.

## Slide 4
Courtesy: Rick Hervig, ASU
Notes: Courtesy of Rick Hervig, SIMS specialist at ASU

## Slide 5
Courtesy: Rick Hervig, ASU
Notes: Courtesy of Rick Hervig, SIMS specialist at ASU

## Slide 6: Most of elements in periodic table are ionized
±10kV
Primary Ions
Cs: +10kV
O−: −13kV
Sample
Secondary Ions
Primary Cs+  Secondary negative ions
Higher electronegativity: O−, C−, S−, Si−
Primary O−  Secondary positive ions
Higher ionization tendency: Mg+, Ca+, Pb+
Modified from Evans Analytical Group: http://www.eaglabs.com/mc/sims-theory.html
Primary Cs+/Sec(−)
Primary O−/Sec(+)
- Noble gases do not ionize.
*Ionization efficiency varies significantly
from 0.1 to 10%. More efficient = better counting statistics = higher precision data.
*Efficiency depends on bonding and structural characteristics of the sample.
Notes: Most elements on the periodic table are ionized by SIMS, but they are ionized with different sources. Positively charged primary ion sources produce negatively charged secondary ions, while negatively charged primary ions yield positively charged secondary ions. An important consideration when using SIMS is the ionization efficiency of the sample. The ionization efficiency is directly related to counting statistics, and therefore the precision that can be achieved. The ionization efficiency itself is influences by the bonding characteristics and the structural characteristics of the samples that you are analyzing.

## Slide 7: Sputtering: Ejection of atoms
0V
-10kV
Primary ions are implanted to the interior of sample.
Primary Ion 133Cs+
Impact energy 20 keV
Ejected atoms
“Sputtering"
amorphous layer
Atoms in the sample are displaced by collision cascade. Lattice structures are destroyed and amorphous layer (~30 nm) is formed.
Some atoms are ejected from the surface = sputtering

## Slide 8: Secondary Ion Generation
0V
-10kV
atomic ions (−)
neutral atoms
molecular ions (−)
neutral molecules
Electrons (−)
e
Extraction
Sample
High Voltage
Most (>90%) atoms and molecules are neutral.
Some are ionized and are accelerated by a 10kV electro-static field
e
e
e
e
e
e
e
e
e
e
e
e
e
e
e
e
e
e
e
e
e
e
e
e
e
e

## Slide 9: Sample charging by the loss of electron
Gold coat
0V
-10kV
atomic ions (−)
neutral atoms
molecular ions (−)
neutral molecules
Electrons (−)
e
Extraction
Sample
High Voltage
Electron is easily ejected out from the sample by the same -10kV field !
e
e
e
e
e
e
e
e
e
e
e
e
e
e
e
e
e
e
e
e
e
e
e
e
e
e
Your samples are insulators!!
e
e
e
e
e
e

## Slide 10: Sample charging by the loss of electron
0V
-10kV
atomic ions (−)
neutral atoms
molecular ions (−)
neutral molecules
Electrons (−)
e
Extraction
Sample
High Voltage
What will happen if electron is lost?
e
e
e
e
e
e
e
e
e
e
e
e
e
e
e
e
e
e
e
e
e
e
e
e
e
e
-9.9kV
-9.5kV
-9.0kV
Secondary ions can not pass through the mass spectrometer, which is tuned for -10kV extraction voltage.

## Slide 11: Sample charging by the loss of electron
0V
-10kV
atomic ions (−)
neutral atoms
molecular ions (−)
neutral molecules
Electrons (−)
e
Extraction
Sample
High Voltage
What would happen if we can not supply any electrons?
e
e
e
e
e
e
e
e
e
e
e
e
e
e
e
e
e
e
e
e
e
e
e
e
e
e
-9.9kV
-9.5kV
-9.0kV
?? kV
Electrostatic discharge!
…on your sample?

## Slide 12: Secondary Ion Mass Spectrometer (SIMS)
±10kV
Secondary Ion
Primary Ion
(+10kV/−13kV)
Magnetic Field
Electrostatic Field
Double focusing
Mass Spectrometer
18O
16O
Ion counters
Sample
E-gun
**Surface analysis**
Simplified Schematic of SIMS
Secondary ions are produced by the sputtering of primary ions
Electron Gun

## Slide 13
……from a certain point of view
From Kevin McKeegan. UCLA SIMS workshop 2009

## Slide 14
WiscSIMS: typically 1 to 40 micron spots, < 1 micron deep
……from a certain point of view
From Kevin McKeegan. UCLA SIMS workshop 2009

## Slide 15
Double Focusing
Mass Spectrometer
Focusing electrons and ions. Tuning the SIMS.
Ion Detection
Sample & Secondary Ions
Electron Gun
Primary Ion
Sources
Total secondary ion travel distance: ~10 m
Notes: If we look at a more detailed schematic of the SIMS what we can see is that there are many many lenses that need to be adjusted in order to do focusing. The flight paths of the primary ion beam, the electrons from the electron gun, and the secondary ions produced all need to be focused well, in order to produce meaningful data. Like the electron microprobe, the focusing is achieved through the use of electrostatic lenses. In addition, there are several places along flight paths where we have slits and apertures, to assist with tuning. These essentially cut out primary and secondary ions. Essentially when tuning up a beam its important to consider how big of an area you can reasonably analyze. Once you know that you can see how well you can focus the beam and see how small you can get it and what kind of beam shape you can achieve. If the beam is too large and/or poorly shaped, you can then introduce slits and apertures to make a smaller beam that’s better shaped.

## Slide 16: Secondary Ion Mass Spectrometer (SIMS)
±10kV
Secondary Ion
Primary Ion
(+10kV/−13kV)
Magnetic Field
Electrostatic Field
Double focusing
Mass Spectrometer
18O
16O
Ion counters
Sample
E-gun
**Surface analysis**
Simplified Schematic of SIMS
Secondary ions are produced by the sputtering of primary ions
Electron Gun
channel plate detector for secondary ion imaging
Notes: So how do we know if we are tuning the beam properly? To help us with this, the SIMS is equipped with a channel plate, which is a detector that can image the secondary ions.

## Slide 17
secondary ions
(a.k.a electron multipliers)
output electrons are converted to an image
fun fact: this is the basis for night vision goggles.
adweb.desy.de
Notes: So what is a channel plate? Essentially it is an array of tiny electron multipliers. When a secondary ion hits the channel plate the signal is converted to an electron, which then cascades through the electron multipler. The electron signal is then processed and converted into an image.

## Slide 18
Beam tuning by imaging of secondary ions
http://epswww.unm.edu/iom/SIMSgear.html
sample holder with Cu grid masking an Al backing
channel plate secondary ion image of 27Al+. Grid spacing is 25 microns.
25 mm
grid
Notes: We can further take advantage of imaging for beam tuning by having a sample with an easy to recognize pattern. Most SIMS labs nowadays have mesh, or grid, tuning samples. Here we can see an example of a SIMS sample holder with a copper grid pressed into an aluminum backing. With the channel plate we can then detect 27Al secondary ions to produce an image. For tuning of the secondary ion flight path, we can optimize the image of the grid with lens alignment that gets rid of tilt and aberration. For other parts of the instrument like primary beam tuning and electron gun tuning we can also take advantage of images projected onto the channel plate, and we can adjust lenses to get the desired image we know correspond to well-focused part of the flight path. In addition we can use the channel plate to see how wide or narrow slits are, and how big or small the sputtered region will be with certain apertures put into place.

## Slide 19: Spot size vs. primary ion beam intensity vs. analytical precision
Cs+ primary beam 10-15 µm spots
Comet sample return; Nakashima et al. (2012)
Cs+ primary beam 1-2 µm spots
Foraminifera; Kozdon et al. (2011)
Primary Ion intensity ~ 2 nA
Precisions of oxygen isotope ratios ~ 0.3‰
3 pA
1-2 ‰
*It is critical to determine how large of a spot size you can accommodate when analyzing regions in samples by SIMS!!
**Often SEM and EPMA are used to determine such sample characteristics.

## Slide 20
Double Focusing
Mass Spectrometer
controlling mass resolution with the entrance slit
Ion Detection
Sample & Secondary Ions
Electron Gun
Primary Ion
Sources
Total secondary ion travel distance: ~10 m
Notes: If we look at a more detailed schematic of the SIMS what we can see is that there are many many lenses that need to be adjusted in order to do focusing. The flight paths of the primary ion beam, the electrons from the electron gun, and the secondary ions produced all need to be focused well, in order to produce meaningful data. Like the electron microprobe, the focusing is achieved through the use of electrostatic lenses. In addition, there are several places along flight paths where we have slits and apertures, to assist with tuning. These essentially cut out primary and secondary ions. Essentially when tuning up a beam its important to consider how big of an area you can reasonably analyze. Once you know that you can see how well you can focus the beam and see how small you can get it and what kind of beam shape you can achieve. If the beam is too large and/or poorly shaped, you can then introduce slits and apertures to make a smaller beam that’s better shaped.

## Slide 21
Mass of 40Ca16O: 55.93494 amu. Mass of 56Fe: 55.95750 amu. ΔM: 0.02256 amu.
Mass resolving power = M/ ΔM. In this case, the minimum mass resolving power to distinguish the signals, 55.95750/0.02256 amu = 2480.
http://epswww.unm.edu/iom/SIMStrace.html
Green: Fe. Blue: CaO
slit
slit
slit
slit
Circular contrast aperture: 750 µm dia.
Wide slit width (> 750 µm). Mass resolving power: 300.
Narrow slit width (~100 µm). Mass resolving power: 2850.
channel plate secondary ion image

## Slide 22
The entrance slit width sets the mass resolving power
BUT…it also limits the amount of secondary ions passing through to the counters, influencing analytical precision. This needs to be considered!
http://epswww.unm.edu/iom/SIMStrace.html

## Slide 23: On our instrument there are actually multiple slits for better tuning control on the mass resolving power
Entrance slit
Field Aperture
Exit Slit
Energy Slit
Notes: Often however, we only have to close one or two slits, because we can achieve good alignment through other focusing mechanisms.

## Slide 24
Double Focusing
Mass Spectrometer
Total secondary ion travel distance: ~10 m
the detection system
Ion Detection
Sample & Secondary Ions
Electron Gun
Primary Ion
Sources
Notes: If we look at a more detailed schematic of the SIMS what we can see is that there are many many lenses that need to be adjusted in order to do focusing. The flight paths of the primary ion beam, the electrons from the electron gun, and the secondary ions produced all need to be focused well, in order to produce meaningful data. Like the electron microprobe, the focusing is achieved through the use of electrostatic lenses. In addition, there are several places along flight paths where we have slits and apertures, to assist with tuning. These essentially cut out primary and secondary ions. Essentially when tuning up a beam its important to consider how big of an area you can reasonably analyze. Once you know that you can see how well you can focus the beam and see how small you can get it and what kind of beam shape you can achieve. If the beam is too large and/or poorly shaped, you can then introduce slits and apertures to make a smaller beam that’s better shaped.

## Slide 25
EM = electron multiplier.
FC = Faraday cup
-detectors are movable.
-selection of EM versus FC depends on the count rate of the signal in question.
Ex: EM’s can become saturated is the count rate of a particular isotope is too high. This will yield unreliable data (and shorten the life of the EM!).
the detection system

## Slide 26
Quantitative analysis of phases, calibrating due to matrix effects.
Molar fraction of Ca-endmember garnet component
diff
measured difference in δ18O, relative to the known value (‰)
Ickert and Stern 2013. Geostandards and Geoanalytical Research

## Slide 27: What causes matrix effects?
-It is an integrated effect, combining:
Influences from the sputtering process;
Transmission into the mass spectrometer;
Transmission through the mass spectrometer;
Influences from the detection system;
The composition of the phase analyzed;
The structure of the phase analyzed.
Essentially, any minor change to a given analytical setup can have a profound influence on the extent of the matrix correction. Therefore, matrix calibration must be performed every time for every session, and must be performed for every phase being analyzed.

## Slide 28
Instrumental bias = (AX/BX)SIMS /(AX/BX)True
O-isotopes (18O/16O) in silicate, carbonates, and oxide minerals: ±10‰
C- isotopes (13C/12C) in calcite: −40‰
Si isotope (30Si/28Si) in quartz: −30‰
matrix effects of common isotope systems measured by the WiscSIMS ims 1280

## Slide 29
How do we know where we are aiming? Surface imperfections can be our friends (just make sure to miss them when analyzing).
optical microscope image of sample surface.
basaltic glass
low-Ca pyroxene
50 µm
corresponding BSE image of sample surface.
Tenner et al. 2009. Chemical Geology
* EPMA/SEM greatly aids with SIMS navigation
Notes: Cracks and pits should be avoided during SIMS analysis because they often have contaminants in them, either from natural processes, or from sample preparation, including epoxy or other mounting materials, and residual grit from polishing.

## Slide 30: Improved optical resolution on the WiscSIMS IMS-1280
Original
(halogen light)
Blue LED
(450 nm)
UV LED
(370 nm)
It is much easier to aim best position for the analysis
Abbe diffraction limit
optical resolution
light source wavelength
numerical aperture
Numbers in images correspond to width and spacing of lines, in microns
Notes: However, even with the UV LED, it is difficult to optically distinguish features less than 1 micron in size.

## Slide 31
Need a 1 micron spot analysis? FIB marking for navigation by SIMS (a.k.a. DIY topography)
Nakashima et al. 2012. EPSL. Particles returned from the comet Wild 2. Note: standards were also FIB marked to account for any topographic effects.
FE-SEM image of FIB-marked square at region of interest. (Depth: ~30 nm)
channel plate secondary ion image of FIB mark
channel plate secondary ion image of SIMS analysis pit
FE-SEM image of SIMS analysis pit.

## Slide 32: Too much topography becomes our enemy…
matrix effect
topography effects
rim
center
Kita et al. 2009. Chemical Geology. Quartz
-due to disturbing electric fields around the sample.
-it is important that samples have flat, untilted surfaces (some micron-scale surface pits and cracks are OK).

## Slide 33
<1 µm of topography from center to edge of grain
analytical precision improves immensely (in this case ± 0.2 ‰ 2SD)
matrix effect
Flat surfaces promote a high degree of analytical precision
Kita et al. 2009. Chemical Geology.

## Slide 34: SIMS yields spatially resolved isotopic data. An in-situ technique
Fish otolith. Each numbered band corresponds to a year of growth (90 = 1990)
500 µm
Matta et al. 2013. Rapid Communications in Mass Spectrometry
misclab.umeoce.maine.edu
Notes: Ootolith = calcifies structure in the inner ear of fish. One annulus per year.

## Slide 35: SIMS yields spatially resolved isotopic data. An in-situ technique
each data point = 10 µm SIMS spot analysis
O- isotope ratios of otoliths are influenced by the temperature and isotopic composition of water. Unaffected by biological processes.
Matta et al. 2013. Rapid Communications in Mass Spectrometry
distance from otolith center (µm)

## Slide 36: SIMS yields spatially resolved isotopic data. An in-situ technique
each data point = 10 µm SIMS spot analysis
Matta et al. 2013. Rapid Communications in Mass Spectrometry
distance from otolith center (µm)
Decreasing water temperature (4 ⁰C for every 1 ‰).
Juvenile: located in shallow, warmer water
Adult: movement into deeper water

## Slide 37: SIMS yields spatially resolved isotopic data. An in-situ technique
each data point = 10 µm SIMS spot analysis
Matta et al. 2013. Rapid Communications in Mass Spectrometry
distance from otolith center (µm)
Decreasing water temperature (4 ⁰C for every 1 ‰).
Saw tooth up: Winter
Saw tooth down: Summer

## Slide 38: SIMS yields spatially resolved isotopic data. An in-situ technique
Decreasing water temperature (4 ⁰C for every 1 ‰).
each data point = 10 µm SIMS spot analysis
Matta et al. 2013. Rapid Communications in Mass Spectrometry
distance from otolith center (µm)
Active seeking of constant water temperatures?

## Slide 39: SIMS yields spatially resolved isotopic data. An in-situ technique
each data point = 10 µm SIMS spot analysis
Matta et al. 2013. Rapid Communications in Mass Spectrometry
distance from otolith center (µm)
rapid growth
no growth
**overall, such data over the life cycle of this fish would not be fathomable by analyzing this otolith by either a bulk technique or a larger spot technique

## Slide 40
‰
‰
‰
decreasing temperature; increasing condensation from vapor.
Why do isotopes make such good proxies for processes? They fractionate. Example: evaporation and condensation effect on oxygen isotope ratios of water vapor at latitude

## Slide 41
Mass difference, relative to 16O: 17O: (17-16)/16 = 1/16
18O: (18-16)/16 = 1/8
δ17O: ‒ 1 ‰
‰
‰
‰
δ17O: ‒ 2.5 ‰
δ17O: ‒ 7.5 ‰
17O fractionates about half as much as 18O

## Slide 42
mass dependent fractionation in terrestrial materials
Notes: On earth we see evidence of mass dependent fractionation in other rocks and waters and when comparing delta 18O to delta 17O, these data define the terrestrial fractionation line, which unsurprisingly has a slope of about ½.

## Slide 43
TFL
mass independent fractionation. slope 1 line
Earth
(light isotope enriched)
Mass dependent fractionation. Slope 0.52.
(heavy isotope enriched)
Mass independent O-isotope fractionation in components from undifferentiated meteorites
Sun
chondrule data: Krot et al. 2009
CAI data: Clayton, 1993
Sun data: McKeegan et al 2011.
Meteorite H2O data: Choi et al. 1998; Sakamoto et al. 2007.

## Slide 44
Protostellar disk surface
Far-UV stellar radiation
Specific CO dissociative wavelengths
C16O
C17O
C18O
1. Equal dissociation of C16,17,18O.
2. C16O dissociative wavelength is dampened. Equal dissociation of C17O and C18O.
3. subsequent cooling
17=18O enriched H2O ice
16O: 17O: 18O:
99.75% 0.038% 0.205%
Mass independent breakdown of CO gas (the dominant species) on the surface of a protostellar disk
Enriched C16O is detected on surfaces of neighboring molecular clouds and protostellar disks (Wilson & Rood, 1994; Sheffer et al., 2002)

## Slide 45
TFL
Earth
(light isotope enriched)
(heavy isotope enriched)
Mass independent O-isotope fractionation in components from undifferentiated meteorites
Sun
chondrule data: Krot et al. 2009
CAI data: Clayton, 1993
Sun data: McKeegan et al 2011.
Meteorite H2O data: Choi et al. 1998; Sakamoto et al. 2007.
mixing of H2O and “Sun” components?
Notes: In contrast materials from chondrites plot along a slope 1 line meaning their ratio of 17 to 18 oxygen in all of these data are 1 to 1. This indicates that chondritic materials fractionated by a mass independent process, completely different than the mass dependent processes that occur on Earth. Because these materials plot along a slope 1 we can define their oxygen isotope ratios in terms of capital delta 17O, which is the per mil deviation in delta 17, relative to the terrestrial fractionation line. Of particular interest is that all chondrite data are enriched in heavy isotopes when compared to the value of the Sun, which, based on its mass can be inferred to represent the bulk oxygen isotope ratio of the solar system itself. Not only that, but theres a wide variability of fractionation, as some measurements of water in chondrite matrices, like this data from sakamor et al 2007, have a capital delta 17 value of +85 per mil, which is over 100 per mil greater than the sun. So overall what kind of mass independent process could have led to these observed values?
You can see that it plots at the light isotope endmember of this fractionation. And the heavy oxygen isotope endmember consists of chondritic water.
In the 1970’s some of the first high precision measurements of extraterrestrial materials were performed by Bob Clayton and Larry Grossman. This was brought mainly by labs gearing up for studying samples returned from the moon, and fortuitously by a meteorite fall in Mexico in 1969 called Allende, which contains large and abundant CAIs. Their mineralogy has been interpreted as the product of primary condensation from extremely hot gas in the early solar nebula. Therefore, the oxygen isotope ratios of minerals in CAIs constrain the oxygen isotope ratio of the early solar nebula itself.

## Slide 46
chondrules are the dominant component of undifferentiated meteorites

## Slide 47
chondrules are the dominant component of undifferentiated meteorites
in between: matrix, which can have materials that interacted with H2O, and CAIs

## Slide 48
ol
lpx
quenched melt
Fe-rich metal
chondrule to chondrule variability in olivine FeO contents, from within the same carbonaceous chondrite
Olivine molar % FeO endmember content: 1.2
Olivine molar % FeO endmember content: 27.7
olivine
quenched melt (SiO2-rich)
Fe-rich metal
olivine: (Mg, Fe)2SiO4
fayalite (FeO endmember): Fe2SiO4
forsterite (MgO endmember) Mg2SiO4
**From olivine-metal phase equilibria, chondrules with FeO-rich olivine formed in more oxidized environments than chondrules with FeO-poor olivine. Was H2O the oxidizing agent?**

## Slide 49
chondrules from a more reduced environment. Less H2O?
chondrules from a more oxidized environment. More H2O?
Heavy isotope-enriched H2O inferred from meteorite matrix phases
**Overall, oxygen isotope ratios of chondrules from carbonaceous chondrites indicates variable mixing of H2O into early Solar System materials**

## Slide 50
Other SIMS techniques: depth profiling
Rick Hervig, UCLA SIMS workshop 2009

## Slide 51
Depth profiling by SIMS is a useful technique for quality control of multi-layered materials
TEM micrograph of a Si-SiGeC multilayer
LeThanh et al. 2002. Materials Science and Engineering

## Slide 52
From Rick Hervig. UCLA SIMS workshop 2009

## Slide 53
SIMS standards on 1 inch diameter SIMS mount
2.5 cm
Another SIMS application: trace analysis (in this case, of hydrogen in upper mantle phases)
Tenner et al. 2009. Chemical Geology

## Slide 54
SIMS phase specific calibrations
Tenner et al. 2012. Contributions to Mineralogy and Petrology

## Slide 55
cpx avg: 0.116 wt.%
glass avg: 4.8 wt.%
glass avg: 4.8 ± 0.6 wt.% H2O
DH20cpx/melt: 0.024 ± 0.004
DH20opx/melt: 0.021± 0.003
High pressure, high temperature experiment 4 GPa, 1405oC
cpx avg: 0.116 ± 0.007 wt.%
opx avg: 0.100 ± 0.004 wt.%

## Slide 56
Combining mineral/melt H2O partition coefficients with mineral modal abundances allows for estimating peridotite/melt H2O partitioning with depth.
Hirschmann et al., 2009. Physics of the Earth and Planetary Interiors

## Slide 57
Combining peridotite/melt H2O partitioning with estimated bulk mantle H2O concentrations can help us determine where hydrous partial melt might be stable in the upper mantle, which might explain observed seismic low velocity zones (LVZ).

## Slide 58
Bulk H2O beneath ridges: 50 – 200 ppm. Not enough to induce hydrous partial melting.
Bulk H2O beneath hotspots and associated w/subduction zones: 300 – 1000 ppm. Enough to induce hydrous partial melting, which may explain seismic low velocity zones at 90 to 220 km depth.

## Slide 59
another SIMS application: rastered isotope imaging
MCF7 cell exposed to 15N-glutamine
Jiang et al. 2014 in press. Methods